Europa 6PLUS

Range

140ºC to 65ºC below ambient

The EUROPAPLUS offers an exceptionally wide operating range in an easy to use portable package. As part of the ISOCAL-6 series the EUROPAPLUS can be used as a Dry Block, a Liquid Bath, a Blackbody Source for infrared thermometers, a Surface Sensor Calibrator and for performance to a few mK (0.001ºC) ITS-90 Fixed Points. The EUROPAPLUS is available in two models, the BASIC (B) and the SITE(S). The B model includes a sophisticated temperature controller with a dual display for Set Temperature and Dry Block Temperature, refer to page 27 for further details. The S model additionally includes an in-built digital indicator to which an external temperature sensor can be connected.

For Dry Block use the unit under test may be compared to the controller value, the use of an external probe connected to the indicator of the Site model which compensates for the temperature gradient and loading errors, is recommended. This gives greater accuracy and traceability meeting the requirements of ISO 9000 and other quality standards.

For liquid bath, surface sensor and blackbody use an external standard probe should always be used, either with the S models indicator or an external indicator. For laboratory use the EUROPAPLUS can be used with a laboratory performance temperature indicator.

Includes as standard: Windows Software, Computer Interface and a Ramp to Set Point Feature.

Increased resolution of ±0.01 available throughout the range via the PC interface and from -19.99 to +99.99 locally on the auto-ranging front display. The controller features multi-point block to display correction giving good absolute accuracy.

The S model has universal sensor input allowing Platinum Resistance Thermometers, Thermocouples (Types K, N, R, S, L, B, PL2, T, J and E) along with Linear Process Inputs including 4-20mA current transmitters to be displayed on the inbuilt indicator. The indicator can be programmed with up to five calibration points to provide high accuracy digital probe matching.

The indicator and controller are both addressable over the communications link.

New! The Site model can now be used with the supplied Cal Notepad software to test thermostats.

Key Features

_ Extreme Low Temperature Operation: -45ºC in an ambient of 20ºC - Use for Triple Point of Mercury, 0ºC, 121ºC (250ºF), 135ºC.

_ Realise ITS-90 Fixed Points: - the EUROPAPLUS can be used with Slim Mercury, Gallium and Triple Point of Water Fixed Point Cells.

_ Calibrate as a Dry or Liquid Bath, thermocouples, platinum resistance thermometers, thermistors, liquid in glass, thermostats, semiconductor types.

_ S Model Includes Universal Input Temperature Indicator allowing for up to five “correction” points to be programmed.

_ Windows Software and PC Interface as Standard.

_ Free Evaluation Report - ask for full data or visit

Options

| 1. Metal Block Insert | 951-02-15 | Standard Insert Included |
| --- | --- | --- |
|  | 951-02-15A | Blank Insert Insert without pockets for local machining |
|  | 951-02-15C | Special Insert Contact AOIP with your requirements |
| Alternative Metal Block Inserts |  |  |
|  | 951-06-07 | Standard Insert type B 13mm , 10mm, 8mm, 5mm and 3.5mm diameter holes, all 157mm deep |
|  | 951-06-08 | Special Insert type C 8mm, 6 x 6.5mm diameter holes, all 157 deep |
| 2. Stirred Liquid Bath | 951-06-01 | Includes a container, magnetic stirrer and probe guide |
| 3. Water/Ice Bath |  |  |
| Thermometer Support Kit | 951-06-03 | Allows three thermometers to be suspended in the bath, including liquid in glass types |
| Oil | 951-06-06 | 0.1L 20ºC to 140ºC |
| 4. Blackbody Target | 951-06-04 | Use with Standard Probe (935-14-82) |
| 5. Surface Sensor Kit | 951-06-02 | Includes an Insert and an angled PRT. |
| 6. Fixed Point Cells | 17724 Slim | Mercury Slim Cell |
|  | D8 | Water Slim Cell |
|  | 17401 Slim | Gallium Slim Cell |
| Standard Probe | 935-14-82 | Platinum Resistance Thermometer |
| UKAS Calibration | UKAS | Calibration available to Order |
| Carrying Case | 931-22-64 | Sturdy case accommodates the unit with room for accessories |

Calibration and Uncertainty

A certificate, traceable to National Standards, is included as standard.

Recommended is an optional UKAS five-point calibration.

The accuracy will depend very much on the mode of use and the types of sensor to be used. Please contact AOIP for tutorials and uncertainty calculations and comprehensive evaluation reports. The EuropaPLUS meets the Calibration Capacity requirements of EA I0/13, “EA Guidelines on the Calibration of Temperature Block Calibrators”.

Model No. EUROPAPLUS

Temperature Range: 65ºC below ambient to +140ºC (absolute minimum -55ºC)

Absolute stability over 30 minutes: Dry Block Bath:               ±0.03ºC

Stirred Liquid Bath:             ±0.025ºC

Ice/Water Bath:                ±0.001ºC

Blackbody Source:              ±0.3ºC

Surface Sensor Calibrator:        ±0.5ºC

ITS-90 Fixed Point Apparatus:    ±0.0005ºC

Computer Interface: Included with Windows Software (page 74)

Cools from: 140ºC to 0ºC in 15 minutes

Heats from: -30ºC to 140ºC in 15 minutes

Best Performance: See Graph

Calibration volume: 35mm diameter by 160mm deep

Standard Insert: 6 pockets, 9.5mm, 8.0mm, 6.4mm,

6.4mm, 4.5mm, 4.5mm, all 157mm deep

Uniformity: ±0.018ºC

Display Resolution: 0.01 -19.99 to 99.99

0.1 -55.0 to -20.0 and 100.0 to 140.0

PC can display 0.01 across whole range with the software included

Indicator units: ºC, ºF, K

Power: 100 to 120V (50 / 60 Hz) or

200 to 240V (50 / 60 Hz)

300 Watts

Overall dimensions: Height 322mm

Width: 176mm

Depth 262mm

Weight: 14kg

How to Order EUROPA-6PLUS

Please specify model type required

Please specify voltage required

Please specify options required